# 1.  Getting Started

This manual contains the information you will need to install and use the EXM-11 interval timer.

EXM-11 provides for timing of intervals between external signals.  The resolution of the timing is 100 ns, with a 100 ppm accuracy.  EXM-11 contains an 8192-entry FIFO memory to allow capture of a series of events happening at a rate faster than they can be read by software.

As shown in the diagram below, the front panel contains BNC connectors for two TTL input signals - gate and signal.  The typical usage of the EXM-11 is that a rising or falling edge (or both) of signal is used as a trigger (i.e., signal is the event being timed), and gate is used as an enable/disable condition (i.e., ignore or trigger on edges on signal).  When a trigger event occurs, the value of a 28-bit counter (with a 10 MHz frequency) is placed in the FIFO, along with the state of the gate and signal inputs.

Control registers and the FIFO reside in the PC I/O space.  Two types of software interfaces are provided: interrupt driven and DMA driven.  In both, the contents of the 32-bit-wide FIFO can be obtained by successive 16-bit reads.  EXM-11 contains no physical jumpers or switches; its configuration (I/O address, interrupt, DMA channel) is programmed using the standard EXM configuration architecture.

## Specifications

The following table defines the power and environmental specifications of the EXM-11.

Characteristic

Value

Temperature

operating

0 to 60°C ambient

storage

-40 to 125°C

Humidity

operating

0 - 90% noncondensing

storage

0 - 95% noncondensing

Altitude

operating

10,000 ft (3000 m)

storage

50,000 ft (15,000 m)

Vibration

operating

0.015 inch (0.38 mm) P-P displacement with 2.5 g peak (max) acceleration over 5-2000 Hz

storage

0.030 inch (0.76 mm) P-P displacement with 5.0 g peak (max) acceleration over 5-2000 Hz

Shock

operating

30 g, 11 ms duration, half-sine shock pulse

storage

50 g, 11 ms duration, half-sine shock pulse

Power

maximum

10 W

typical

7.5 W

Current

maximum

5V @ 2.0A

typical

5V @ 1.5A

Weight

4 oz (120 g)