4.  Programming Interface

This chapter contains information needed to write software drivers for the EXM-11.

The EXM-11 defines the following registers in the I/O space.

Bit 7

Bit 6

Bit 5

Bit 4

Bit 3

Bit 2

Bit 1

Bit 0

I/O port

Device ID Register

1

1

1

1

0

1

0

0

100

Config Option Byte 1 Reg

DMA channel

Interrupt

I/O base

Cden

102

FIFO Read Register

XX0

Status/Control Register

OVR

Dual

SEN

SPL

GEN

GPL

ACQ

ENF

XX2

Software Trigger/Test Reg

XX4

The I/O port denoted XX0 represents the selected I/O base address for the EXM-11, XX2 represents this address plus two, and so on.

The first two are standard EXM registers for device identification and configuration.  The EXM-11 responds to accesses to these ports only if EXMbus line -EXMID is asserted.  An 8-bit read from I/O address 100h returns the value F4, the device ID of the EXM-11.  A read/write configuration register appears at I/O address 102h.  The bit encodings are defined in the previous chapter.

DMA channel selects the DMA channel to be used, or specifies that no DMA channel will be used.

Interrupt selects the interrupt to be generated, or specifies that no interrupt is to be generated.

I/O base  selects the place in the I/O space in which the I/O port area of the EXM-11 appears.

CDEN specifies whether the EXM is disabled or enabled.  If disabled, the EXM will not respond to any I/O or memory addresses and will not assert an interrupt output; it will only respond to reads from I/O port 100h and reads and writes from I/O port 102h, and then only if EXMbus line -EXMID is asserted.

The FIFO read register is an 8/16-bit read-only register.  It is used to read the oldest 32-bit value in the FIFO, either by a DMA channel (via two 16-bit accesses) or by software (via two 16-bit accesses or four 8-bit accesses).

Actual values in the FIFO have the following representation

31

29

28

27

0

G

S

TIME

but the 32-bit value read has the following representation

31

30

29

28

27

0

NF

NE

G

S

TIME

NF	Not Full.  If 1, the FIFO was not full at the time that the byte containing the NF flag was read from the FIFO read register.  When the FIFO is read with a DMA transfer, this bit is masked and always returned as 0.

NE	Not empty.  If 1, the FIFO was not empty at the time that the byte containing the NE flag was read from the FIFO read register.  If 0, none of the other 32 bits have meaning.  The preceding applies when the FIFO is read with a nonDMA transfer; when the FIFO is read with a DMA transfer, this bit is masked and always returned as 0.

G	A function of the gate front-panel input at the time this word was stored in the FIFO.  This is explained in more detail later.

S	A function of the signal front-panel input at the time this word was stored in the FIFO.  This is explained in more detail later.

TIME	The value of the 28-bit counter at the time this word was stored in the FIFO.

If the DMA channel field of the configuration register specifies one of the two DMA channels or specifies DMA disabled and 16-bit FIFO transfers, the FIFO read register behaves as a 16-bit I/O port.  Each pair of reads returns a word from the FIFO and removes that word from the FIFO.  The first read returns the low-order 16 bits and the second read returns the high-order 16 bits.

If the DMA channel field specifies DMA disabled, 8-bit FIFO transfers (the encoding 00), the FIFO read register behaves as an 8-bit I/O port.  Each set of four reads returns a word from the FIFO and removes that word from the FIFO.  The first read returns bits 0-7, the second returns bits 8-15, the third returns bits 16-23, and the fourth returns bits 24-31.

For instance, the following C language code segment reads a word from the FIFO assuming the EXM-11 has been programmed to be in nonDMA 16-bit transfer mode.

low  = inpw(EXM11+0);

high = inpw(EXM11+0);

if ((high & 0x4000) == 0) 	/* if empty ... */

if ((high & 0x8000) == 0) ...	/* if full ... */

The bits of the status/control register are defined below.

OVR	Overrun.  If set, an overrun has occured (an entry could not be placed in the FIFO because it was full).

DUAL	Dual-edge trigger.  If set, both edges of signal create a trigger event.  If 0, only the leading edge of signal causes a trigger.

SPL	signal polarity.  If set, signal is positive true, and its leading edge is defined as the transition from low voltage to high.  If 0, signal is negative true, and its leading edge is defined as the transition from high to low.

SEN	signal enable.  If set, signal is sampled for trigger events.  If 0, transitions on signal do not cause trigger events.

GEN	gate enable.  If set, gate is used to qualify signal trigger events.  If 0, signal trigger events occur regardless of gate.

GPL	gate polarity.  If set, gate is positive true.  If 0, gate is negative true.

ACQ	Acquire data.  If set, trigger events are captured in the FIFO.  If 0, they are not.

Clearing this bit also clears OVR.  To clear an overflow condition, one should make the FIFO exit the full state before changing ACQ back from a 0 to 1 (or else setting ACQ will set OVR again).

ENF	Enable FIFO.  If set, the FIFO operates normally.  If cleared, the FIFO is cleared.  Clearing the FIFO does not affect the 28-bit counter.  (The counter is cleared only by a hardware reset.)

The software-trigger is an 8-bit read-only test register.  Reading it causes a trigger event.  The data returned from the read is not meaningful.

## Trigger Events

A trigger event occurs and causes an entry to be placed in the FIFO memory when the ENF control bit is set and one of the following conditions is present:

•	The signal input causes the trigger event.  For this to occur, the SEN and ACQ bits must be set and either the GEN bit must be clear or the gate input must be asserted.  When these conditions exist, a trigger occurs on either edge of signal if DUAL is set or on the leading edge of signal if dual is clear (0).

•	The ACQ bit is set and either edge of the gate input is detected.

•	A read occurs to the test register.

These are summarized in the table on the following page.  This table also shows the values stored in bits 28 and 29 of the FIFO entry when a trigger event occurs.

gate and signal are sampled at 10 MHz intervals.  An edge is defined as a logical inversion from one sample to the next.  When the gate input is used as a qualifier on signal, gate needs to obey setup and hold times of 100 ns with respect to the edges of signal.

The NE (not empty) flag returned from reads of the FIFO will not be asserted until at least two trigger events have occurred after reset.  It will become unasserted (indicating empty) when there are zero or one entries left in the FIFO.

## Interrupts and DMA Operation

If an interrupt is designated in the configuration register, the interrupt is signalled in the following two circumstances:

•	DMA is not enabled (configuration register bit 7 is 0) and the FIFO is "half full" (the FIFO currently contains 4097 or more entries).

•	DMA is enabled and the TC (terminal count) signal is asserted on the current DMA transfer.  The interrupt is deasserted upon the next read to the status/control register.  The EXM-11 assumes that the TC signal only occurs during a DMA transfer for the high-order 16 bits of the FIFO; the results are unpredictable if this is not the case.

DMA operation of the EXM-11 is enabled by specifying DMA channel 5 or 6 in the configuration register.  In this case, the EXM-11 requests a DMA transfer if the FIFO is not empty.  DMA requests continue to be asserted until a read of the high-order 16 bits of the FIFO occurs and the FIFO becomes empty.

In the PC architecture, DMA channels 5 and 6 are 16-bit DMA channels.  The EXM-11 does not support 8-bit DMA transfers.

The following code example shows how to establish an interrupt handler for IRQ11.  The language is C and the operating system is assumed to be MS-DOS.

outp(EXM11 + 2, 0);	/* clear fifo */

outp(EXM11 + 2, 0x03);	/* enable acquisition */

_dos_setvect(0x73,myhandler);	/* install IRQ11 handler */

outp(0xa1,inportb(0xa1) & 0xf7);	/* enable IRQ11 */

...

void interrupt myhandler() {

outp(0xa0,0x20);

outp(0x20,0x20);

...

The following sets up DMA channel 5 and interrupt IRQ15.

_dos_setvect(0x77,myhandler);

outp(0xa1,inportb(0xa1) & 0x7f);

outp(0xd4,0x05);	/* set channel 5 mask bit */

outp(0xd6,0x05);	/* set up mode register */

outp(0xd8,0);	/* clear byte point FF */

outp(0xc4,buffer_adr_low);	/* load buffer address */

outp(0xc4,buffer_adr_mid);

outp(0xc6,count_low);	/* load count */

outp(0xc6,count_hi);

outp(0x8b,buffer_adr_hi);

outp(0xd6,0x01);	/* enable channel 5 */

/* Now wait for IRQ15 */

If your processor is an EPC-4 with a BIOS level prior to 2.18 or an EPC-5 with a BIOS level prior to 3.03, the BIOS contains an error inhibiting 16-bit DMA.  To correct this, insert the following two lines immediately after the the third line above (the line setting the mask bit).

outp(0x26,0x0845);

outp(0x24,0x05);

The following code segment, which is not specific to the EXM-11, shows how to compute the address for the DMA controller.

farptr = (char far *) &buffer;

dma_adr = ((((unsigned long) farptr >> 16) << 4) + ((unsigned long) farptr &

0xffff)) >> 1;

buffer_adr_lo = dma_adr & 0xff;

buffer_adr_mid = (dma_adr >> 8) & 0xff;

buffer_adr_hi = (dma_adr >> 15) & 0xff;