# 1.  Product Description

This manual contains the information required to install and use the EXM-21 digital-to-analog converter expansion module.

The EXM-21 is a dual, 12-bit digital-to-analog (D/A) EXM expansion interface card. It contains an EXM expansion interface, two 12-bit D/A converters, output amplifiers, and a reference generator.  The output stages of the EXM-21 provide unipolar or bipolar outputs up to ±10V, and a 4-20mA current loop. A DC-DC converter is provided for +5V only operation.

The EXM-21 is software- and output-compatible with the Keithley Metrabyte DAC-02 card (with the exception of calibration procedures -- refer to Chapter 3). The EXM-21 contains a female DB-25 connector on the face plate, along with four trimpots accessible through the front panel.

## µ §

## Specifications

The following tables define the specifications of the EXM-21.

| Characteristic |  | Value |
| --- | --- | --- |
| Environmental |  |  |
| Temperature | operating | 0° to 60° C (* see below) |
|  | storage | -40° to 125° C |
| Humidity | operating | 0 to 95% (non-condensing) |
|  | storage | 0 to 95% (non-condensing) |
| Vibration | operating | .015"PP 2.5g (max) 5 to 2000 Hz |
|  | storage | .030"PP 5g (max) 5 to 2000 Hz |
| Shock | operating | 30g 11 msec duration |
|  | storage | 50g 11 msec duration |
| Electrical |  |  |
| Current | maximum | +5V @ 600 mA |
|  | typical | +5V @ 400 mA |
| 10V Reference Generator | initial accuracy | ± 10.5 mV |
|  | temperature coefficient | 3 ppm typical, 50 ppm max |
|  | LT stability | 15 ppm typical |
| 5V Reference Generator | initial accuracy | ± 7.5 mV |
|  | temperature coefficient | 3 ppm typical, 50 ppm max |
|  | LT stability | 15 ppm typical |
| Fuse | External +5V | Schurter OMF 2A (socketed) |

Table 1. EXM-21 Environmental and Electrical Specifications.

* Upper temperature limit degrades 2° C per 1000 ft. elevation.  Maximum  
   elevation 10,000 ft.

Refer to Table 2 below for output range specifications.

| Description |  | Specification |
| --- | --- | --- |
| System |  |  |
| Channels |  | 2 |
| Resolution |  | 12 bit (1 part in 4096) |
| I/O Address |  | Software-configured on any 8-bit boundary |
| Relative Accuracy |  | 1/2 LSB (0.01% of output voltage) maximum |
| Current Loop |  | 4 mA to 20 mA |
| Differential Linearity |  | 1/2 LSB maximum |
| Output Ranges when using Internal -5V Reference |  | Unipolar: 0V to +5V Bipolar: -5V to +5V |
| Output Ranges when using Internal -10V Reference |  | Unipolar: 0V to +10V Bipolar: -10V to +10V |
| Reference |  |  |
| DC Reference Input Range |  | 0 maximum to -10V minimum |
| Variable Reference Input Range |  | ± 10V maximum |
| Reference input resistance |  | 7 K Ohm minimum 11 K Ohm typical 20 K Ohm maximum |
| Output |  |  |
| Output Amplifiers Offset | unipolar bipolar | < 50 uV unipolar @ 25 °C < 120 uV unipolar 0° to 60 °C < 100 uV bipolar @ 25 °C < 240 uV bipolar 0° to 60 °C |
| Voltage output impedance |  | 0.1 ohm maximum |
| Voltage output drive current |  | ± 5 mA minimum |
| Current Loop |  |  |
| Current Loop Compliance Voltage |  | 8V to 36V |
| AC Reference |  |  |
| AC Reference Operation |  | ± 10V, 2-quadrant and 4-quadrant |
| Feed through |  | 5 mV @ 10 KHz, ± 5V Reference typical |
| Accuracy/Frequency |  | 12 bit (0-1 KHz Reference) |

Table 2. Output Specifications

NOTES