# 3. Theory of Operation

Basics of EXM-21 Operation

The EXM-21 uses Analog Devices AD7548 D/A converters (or compatible).  The inputs to the A/D converter are a negative reference voltage (Vref) and a digital value between 0 and 4096. The A/D converter outputs a voltage  between 0 volts and -(Vref) volts, proportional to the digital value programmed.

Typically, a DC voltage is used as a reference to generate a single DC output voltage for each digital value programmed into the D/A converter.  An AC reference may also be used to generate a variable output voltage for each digital value.  One use of this is in synchro-digital and resolver applications for sine/cosine generation.  Specific applications of an AC reference are beyond the scope of this manual.

µ §The amplifier  section of the EXM-21 buffers the output of the D/A converter and provides conversions to the bipolar or current loop outputs.  When using an AC Reference, the outputs are termed 2-Quadrant (unipolar output) and 4-Quadrant (bipolar output).  Figure 3 shows that a 2-Quadrant signal is made by varying the reference voltage between positive and negative voltages, and by keeping the programmed value positive.  When using the 4-Quadrant (Bipolar) output, the programmed value can be programmed to be a negative value as well.

Figure 3. Two-quadrant and Four-quadrant operation.

Each D/A converter of the EXM-21 is calibrated separately, and must be re-calibrated for each new reference source or reference voltage used.

External Connections to the EXM-21

Front Panel Connector

The front panel of the EXM-21 has a DB-25 connector which is compatible with the Keithly Metrabyte DAC-02 connector.

| Pin | Signal | Pin | Signal |
| --- | --- | --- | --- |
| 1 | Digital ground | 14 | -10V Reference |
| 2 | Analog ground | 15 | -5V Reference |
| 3 | Analog ground | 16 | D/A #1 Ref Input |
| 4 | Analog ground | 17 | D/A #1 Bipolar Output |
| 5 | Analog ground | 18 | D/A #1 Unipolar Output |
| 6 | Analog ground | 19 | D/A #1 4-20 mA Output |
| 7 | Analog ground | 20 | -10V Reference |
| 8 | Analog ground | 21 | -5V Reference |
| 9 | Analog ground | 22 | D/A #0 Reference Input |
| 10 | Analog ground | 23 | D/A #0 Bipolar Output |
| 11 | Analog ground | 24 | D/A #0 Unipolar Output |
| 12 | Analog ground | 25 | D/A #0 4-20 mA Output |
| 13 | Fused +5V |  |  |

Table 5. DB-25 Pin-out

Reference Voltages

The Reference voltage for each D/A converter is supplied through the front panel connector and may be connected to one of the internal references supplied on pins 14 & 15 and pins 20 & 21, or to an external reference source.

When using the current loop output, the internal -5V reference must be used.

The internal reference generators produce stable -5V and -10V reference voltages from a laser-trimmed -10V source.

Analog Outputs

Three types of analog outputs are provided for each D/A converter: unipolar, bipolar and  current loop.  The unipolar output has a range between 0V and -(Vref).  The bipolar output has a range between (Vref) and -(Vref).  The current loop has a range between 4mA and 20mA.

When using the current loop output, an external voltage source must be supplied to source the current.  Two typical current loop configurations are shown in Figure 4.  In each configuration, the current is measured by measuring the voltage across the resistor R.  Note that when using the current loop output, a minimum of 8 volts relative to the EXM-21 analog ground must be present at the current loop output pin at all times.

µ §

Figure 4. Typical Current Loop Configurations.

When using an AC reference,  a 2-Quadrant output may be taken from the unipolar output, or a 4-Quadrant output may be taken from the bipolar output.

The following tables summarize all output ranges which use the internal EXM-21 references, as well as formulas for external reference voltages.

|  |  |  |  |  | Output |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| D/A | Output Mode | Out-put Pin | Int Vref | Int Vref Conn | Min | Max | Step Size (1LSB) |
| D/A #0 | Unipolar 0-5V | 24 | -5V | 21-22 | 0.00V | 4.99878V | 1.221 mV |
| D/A #1 | Unipolar 0-5V | 18 | -5V | 15-16 | 0.00V | 4.99878V | 1.221 mV |
| D/A #0 | Unipolar 0-10V | 24 | -10V | 20-22 | 0.00V | 9.99756V | 2.441 mV |
| D/A #1 | Unipolar 0-10V | 18 | -10V | 14-16 | 0.00V | 9.99756V | 2.441 mV |
| D/A #0 | Bipolar ± 5V | 23 | -5V | 21-22 | -4.99756V | 5V | 2.441 mV |
| D/A #1 | Bipolar ± 5V | 17 | -5V | 15-16 | -4.99756V | 5V | 2.441 mV |
| D/A #0 | Bipolar ± 10V | 23 | -10V | 20-22 | -9.99512V | 10V | 4.883 mV |
| D/A #1 | Bipolar ± 10V | 17 | -10V | 14-16 | -9.99512V | 10V | 4.883 mV |
| D/A #0 | Current Loop | 25 | -5V | 21-22 | 4 mA | 19.996 mA | 3.906 mA |
| D/A #1 | Current Loop | 19 | -5V | 15-16 | 4 mA | 19.996 mA | 3.906 mA |

Table 6. Output Ranges for Internal References.

| Output Mode | Vref | Min | Max | Step Size (1LSB) |
| --- | --- | --- | --- | --- |
| Unipolar | Ext | 0.00V (when DAC value = 0) | -Vref * (4095 / 4096) (when DAC value = 4095) | -Vref /4096 |
| Bipolar | Ext | Vref * (2047/ 2048) (when DAC value = 4095) | -Vref (when DAC value = 0) | -Vref /2048 |
| 2 Quadrant | AC | 0.00V | -Vref * (4095 / 4096) | - |
| 4 Quadrant | AC | Vref * (2047/2048) | -Vref | - |

Table 7. Output Ranges for External References.

EXM-21 Programming Interface

Registers

The EXM-21 is programmable at the register level using I/O (Input/Output) instructions.

Data format for the D/A registers is as follows:

|  | Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Low Byte (Least Significant Byte): |  |  |  |  |  |  |  |  |
| Base + 0 or 2 | B3 | B2 | B1 | B0 | x | x | x | x |
| High Byte (Most Significant Byte): |  |  |  |  |  |  |  |  |
| Base + 1 or 3 | B11 | B10 | B9 | B8 | B7 | B6 | B5 | B4 |

Table 8. D/A Registers Data Format

Registers Base+0 and Base+1 are for D/A #0 and registers Base+2 and Base +3 are for D/A #1.  These registers are write-only and cannot be read.  The low byte is written first (Base+0  or Base+2  register) followed by the high byte (Base+1 or Base+3  register).  When the low byte is written, the data is stored in an intermediate register.  When the high byte is written, the entire 12 bit word is loaded into the D/A converter and the output voltage changes.

Note that for unipolar and current loop outputs, the output voltage or current is proportional to the programmed value.  However, for the bipolar output, the voltage is inversely proportional to the programmed value.  For example, if a programmed value of 0 is written to the D/A converter, the resulting bipolar output voltage would be the maximum positive voltage.  If 2048 is written, the resulting bipolar output voltage would be 0.00 volts.  If 4095 is written, the resulting bipolar output voltage would be the maximum negative voltage.

Programming Example

The following programming example is written in Microsoft C.  The only command which is dependent on the Microsoft version is the outp command which writes a byte to an I/O port. The syntax of this command is:

int outp( unsigned port_adrs, int value);

Descriptions of each function of the example are included in the comments before the functions.

/*

* Function:    Set_DAC

*

* Description: Writes a 12bit value to the D/A converter specified.

*

* Parameters:  exm21_base      Base address of the EXM-21

*              dac             EXM-21 D/A Converter to use (0 or 1)

*              count           Digital value to write (0 to 4095)

*

*/

void

Set_DAC( unsigned int exm21_base,

unsigned short dac,

unsigned short count )

{

outp((exm21_base + 2 * dac), (count << 4));

outp((exm21_base + 2 * dac + 1), (count >> 4));

}

/*

* Function:    Set_DAC_Volts

*

* Description: Writes a value to the D/A converter specified to

*              generate a given voltage.

*

* Parameters:  exm21_base      Base address of the EXM-21

*              dac             EXM-21 D/A Converter to use (0 or 1)

*              volts           Output voltage specified (in volts).

*              vref            Reference voltage connected to

*                              D/A converter.

*              unipolar        TRUE if Unipolar output,

*                              FALSE if Bipolar.

*/

void

Set_DAC_Volts( unsigned int exm21_base,

unsigned short dac,

double volts,

double vref,

unsigned short unipolar )

{

unsigned short count;

if( unipolar ){

count = -(double) 4096 * volts / vref;

Set_DAC( exm21_base, dac, count);

} else {

count = -(double) 2048 *( (double)1 + volts/vref);

Set_DAC( exm21_base, dac, count);

}

}

/*

* Function:    Set_DAC_CL

*

* Description: Writes a value to the D/A converter specified to

*              generate a given current at the Current Loop output.

*

* Parameters:  exm21_base      Base address of the EXM-21

*              dac             EXM-21 D/A Converter to use (0 or 1)

*              mamps           Output current specified (in mA)

*                              (4mA to 20mA)

*/

void

Set_DAC_CL(    unsigned int exm21_base,

unsigned short dac,

double mamps )

{

unsigned short count;

count = (double) 256 * mamps - (double)1024;

Set_DAC( exm21_base, dac, count);

}

Calibration

Calibration of the EXM-21 should be performed periodically to maintain the highest accuracy, or whenever a new reference voltage or reference source is used.  The EXM-21 is calibrated at the factory using the internal -10V reference for the unipolar and bipolar outputs.

For laboratory environments, the EXM-21 should be re-calibrated at a six-month to one year interval.  For more rigorous conditions where extremes of temperature, vibration or humidity are encountered, a three-month interval is recommended.

The location of the trimpots used during calibration is shown in Figure 5.

Figure 5. Trimpot Locations

Calibration Procedure

Note: This calibration procedure assumes the use of and familiarity with MS-DOS commands, especially the DEBUG utility provided with MS-DOS.  If this is not available, please contact RadiSys Customer Service for additional MS-DOS software which may be used to calibrate the EXM-21.

Equipment Required:

·	RadiSys EMC or EPC system including CPU, video adapter, monitor, keyboard and disk drive

·	Non-metallic slotted calibration tool

·	Male DB-25 connector with connections to internal or external references as required by the specific application. For AC reference applications, the AC source should be used with the input set to its peak output voltage that will be used. (See Selecting Output Ranges, page 14)

·	Digital Volt Meter (DVM) and Digital Current Meter (DCM) accurate to 5 digits

·	Voltage source capable of sourcing up to 10mA at 8V or greater.

·	MS-DOS and DOS DEBUG utility.

Procedure:

Vref0 = Reference voltage for DAC#0

Vref1 = Reference voltage for DAC#1

1.	Install the EXM-21 into the RadiSys computer system and configure the BIOS setup screen. (See Chapter 2).

(All DEBUG programming examples will assume the EXM-21 base address is at 248h.)

Install the Male DB-25 connector with application references attached.

2.	Power up the RadiSys computer system and boot to DOS. At the DOS prompt, type DEBUG.

3. 	Program DAC #0 and DAC #1 to 4095. In debug, type:

-o 248 ff

-o 249 ff

-o 25a ff

-o 25b ff

4.	Connect the DVM (+) lead to the DB-25 at pin 24.  Connect the DVM (-) lead to the DB-25 at pin 2.

5.	Calculate the calibration voltage for channel 0 as follows;  write the result in the space provided.

V0cal	=  -Vref0 * 0.999756	= _____________

6. 	Adjust FSAJ0 on the EXM-21 front panel until the DVM matches the value calculated in step 5.

7. 	Connect the DVM (+) lead to the DB-25 at pin 18. Connect the DVM (-) lead to the DB-25 at pin 2.

8.	Calculate the calibration voltage for channel 1 as follows;  write the result in the space provided.

V1cal	=  -Vref1 * 0.999756	= _____________

9. 	Adjust FSAJ1 on the EXM-21 front panel until the DVM matches the value calculated in step 8.

If your application is not using current loop outputs, the EXM-21 is now calibrated for your application.  To use the current loop outputs, the -5V internal reference must be used. Continue with step 10.

10.	Program DAC #0 and DAC #1 to 0. In debug, type:

-o 248 0

-o 249 0

-o 25a 0

-o 25b 0

11.	Connect the DCM (+) lead to the Voltage source output.

Connect the DCM (-) lead to the DB-25 at pin 25.

Connect the Voltage source ground to the DB-25 at pin 2.

12. Adjust CLAJ0 on the EXM-21 front panel until the DVM reads 4.000 mA.

13. Connect the DCM (+) lead to the Voltage source output.

Connect the DCM (-) lead to the DB-25 at pin 19.

Connect the Voltage source ground to the DB-25 at pin 2.

14. Adjust CLAJ1 on the EXM-21 front panel until the DVM reads 4.000 mA.

The EXM-21 is now calibrated.